## Slide 1: MIDAS Short Course
Stefan Ritt (Paul Scherrer Institute, Switzerland)
Pierre André Amaudruz (TRIUMF, Canada)

## Slide 2: MIDAS short course
What this is all about
  Data acquisition (DAQ) in nuclear and particle physics
  Basic principles of MIDAS DAQ system
  Usage of the system
  Tools, examples
  Online demo
  Goal: Be able to install, configure and run MIDAS
Whom
  Experimentalists with some computer knowledge
  Test set-ups for R&D
  Small and medium size experiments

## Slide 3: A few remarks
This is a course!
  if something is unclear, please ask immediately
  if you have more general questions, keep them until the end of each section
  Coffee break upon request
  On-line demo is flexible, ask for certain things
Slides and example programs are on CD and available online at http://midas.psi.ch/cgi-bin/cvsweb/midas/doc/course/
Basic C and data analysis knowledge is assumed
Example programs partially simplified, miss status checking
Authors are available during the whole conference

## Slide 4: The Authors
1978 degree in electr. engineering at the Technical Institute, Geneve, Switzerland
1979-1981 BBC, Switzerland
1981-1988 SIN (PSI), MWPCs hardware, DAQ (PDP11+VAX)
1988-1993 CHAOS project at TRIUMF
since 1993 staff member of DAQ group at TRIUMF
1993 Ph.D. in Physics at the University of Karlsruhe, Germany
1993-2000 postdoc at University of Virginia, project manager of  experiment
since 2000 research scientist (staff) at PSI, technical coordinator of MEG experiment
Other interests: Chip design, control systems
Stefan Ritt
Pierre André Amaudruz

## Slide 5: Paul Scherrer Institute
590 MeV Protons
Synchrotron light source SLS
Energy research
Proton irradiation

## Slide 6: Agenda
Introduction
Internal structure
  Remote Procedure Calls
  Buffer Manager
  Online Database
  Slow Control System
  History System
  Alarm System
Frameworks
  Frontend
  Analyzer
Online demo
  Installation
  Configuration
  Running a small experiment
Utilities and usage
Known deployment
Advanced features
  Customization of run control and monitoring
  Special front-end features
Advanced online demo
Part I
Part II

## Slide 7: Why MIDAS?
Fulfill needs of certain experiments (~thousand channels, ~MB data rate)
Incorporated slow control system
Integrated data analysis (histogramming)
Operating system independent
Hardware independent
Quick installation
Flexible, useful for other experiments
Easy application programming
Free (GPL)

## Slide 8: What is MIDAS?
DAQ Framework
  Event transport (multiple producers & consumers)
  Central configuration
  Frontend and analyzer framework
  You have to write “user code”
C library + applications
Contains logging, slow control, alarm system, electronic logbook: Maximum Integrated Data Acquisition System
Suitable for small tests/experiments to medium size experiments, not (?) for HEP experiments
Not to be confused with ESO-Midas, Midas/UK, Auto service… Let’s call it “MIDAS/PSI”

## Slide 9: What is MIDAS?
~100.000 lines of C code
Optimized for event-based DAQ, typically nuclear and particle physics
Contains driver library for CAMAC, VME, FASTBUS
MIDAS is free and under the GPL
Two “main” developers (S. Ritt and P. Amaudruz)
Support by PSI and TRIUMF
http://midas.triumf.ca
amaudruz@triumf.ca
http://midas.psi.ch
Stefan.ritt@psi.ch

## Slide 10: Small (test) system
CAMAC interface
(HYTEC $3000)
PC ($1000)
Linux / Windows
ROOT / PAW
1MB/sec

## Slide 11:  Experiment @ PSI
Front-end PC
VME-Crate
Fastbus-Crate
CAMAC crates

## Slide 12: Medium size experiment
Monitoring
On-line
Analysis
Control
Logging
DLT
VME CPU
VME CPU
Fastbus
Gigabit-Ethernet ~tens of MB/sec
CAMAC
12345
12345
HV
Slow control
Beam line

## Slide 13: Limits of MIDAS
Front-end
Front-end
Front-end
Front-end
Front-end
Front-end
Front-end
Front-end
Analysis
Analysis
Analysis
Analysis
Analysis
Analysis
Analysis
Analysis
Switch
MIDAS not optimized for distributed event building
Better systems: CODA (coda.jlab.org)
“MIDAS is not for LHC experiments, but to develop them”

## Slide 14: History
1990
1992
1994
1996
1998
2000
2002
HIX-System
Oct. ’93 MIDAS start
Implementation
 Experiment
  e 
MIDAS becomes standard DAQ
at PSI and TRIUMF
TRIUMF
CHAOS
ISAC
TWIST
Real-Time ‘97
PSI

## Slide 15: Remote Procedure Calls

## Slide 16: Platform independence
Operating system
independent calls
System call layer
Buffer Manager
bm_open_buffer()
bm_send_event()
...
Online Database
db_find_key()
db_set_value()
...
ss_open_shared_memory()
conditional compiling
Windows NT
OpenFileMapping()
VMS
ppl$create_
shared_memory()
UNIX
shmget()
93%
7%

## Slide 17: MIDAS RPCs
Computer A
program
subroutine
call
program
continues
subroutine

## Slide 18: MIDAS RPCs
subroutine
Computer A
program
subroutine
call
Computer B
subroutine
stub
dispatcher
program
continues
send result
receive
result

## Slide 19: MIDAS RPCs
main()
{
…
sub1(a, &b);
…
}
sub1(int a, int *b)
{
*b = 2*a;
}
No IDL, stubs are produced “by hand”
sub1(int a, int *b)
{
rpc_call(ID_SUB1,a,b);
}
rpc_dispatch(id,…)
{
if (id==ID_SUB1)
sub1(a,b);
}
TCP/IP
sockets
Computer A
Computer B

## Slide 20: MIDAS server
server parent
process
mserver
or inetd
Remote
Client
connect
port 1175
Each client connected to own subprocess
Access control through (x)inetd and optional mserver password
RPC layer completely transparent -> exactly same code if running local or remotely
Little/big endian negotiation/conversion
Motorola: big-endian
Intel: little-endian
0x12
0x34
0x34
0x12
= 0x1234
adr0
adr1
spawned
subprocess
mserver
RPC calls

## Slide 21: Why not standard RPCs?
Availability (VxWorks, RT-Linux, …)
Performance
  RPC layer is speed optimized
  Data representation is negotiated
  Some functions like send_event can bypass part of the standard RPC-layer for buffered and asynchronous RPC
  -> 11MB/sec over 100MBit Ethernet
Dynamic extensions at run-time
Asynchronous notification (-> callbacks)
No IDL compiler necessary, simple syntax

## Slide 22: User-level RPCs
#include "midas.h"
#include <stdio.h>
#define RPC_MYTEST 1
RPC_LIST rpc_list[] = {
{ RPC_MYTEST, "rpc_mytest",
{{TID_INT, RPC_IN},
{TID_FLOAT, RPC_IN},
{TID_INT, RPC_OUT},
{TID_FLOAT, RPC_OUT},
{0} }},
{ 0 }
};
main()
{
char host_name[80];
HNDLE hConn;
INT status;
INT i;
float f;
printf("Remote host name: ");
ss_gets(host_name, sizeof(host_name));
/* register this as an RPC client with rpc_list */
rpc_register_client("MYCLIENT", rpc_list);
/* connect to RPC server */
rpc_client_connect(host_name, 1750, "", &hConn);
/* rpc_mytest just doubles numbers */
rpc_client_call(hConn, RPC_MYTEST, 1, 2, &i, &f);
printf("\nResult should be: 2 4.0\n");
printf("Result is: %d %1.1f\n", i, f);
/* disconnect this client from server */
rpc_client_disconnect(hConn, FALSE);
return 1;
}
#include "midas.h"
#include <stdio.h>
#define RPC_MYTEST 1
RPC_LIST rpc_list[] = {
{ RPC_MYTEST, "rpc_mytest",
{{TID_INT, RPC_IN},
{TID_FLOAT, RPC_IN},
{TID_INT, RPC_OUT},
{TID_FLOAT, RPC_OUT},
{0} }},
{ 0 }
};
INT rpc_mytest(INT i, float f, INT *i1, float *f1)
{
*i1 = i*2; *f1 = f*2;
return 1;
}
INT rpc_dispatch(INT index, void *prpc_param[])
{
INT status;
if (index == RPC_MYTEST)
return rpc_mytest(CINT(0), CFLOAT(1), CPINT(2), CPFLOAT(3));
return RPC_INVALID_ID;
}
main()
{
rpc_register_server(ST_REMOTE, "", 1750, NULL);
/* Register function list. Calls get forwarded to rpc_dispatch */
rpc_register_functions(rpc_list, rpc_dispatch);
/* Server loop */
while (cm_yield(1000) != RPC_SHUTDOWN);
rpc_server_shutdown();
return 1;
}
Server
Client

## Slide 23: The MIDAS Buffer Manager

## Slide 24: Buffer Manager
Data
Producer
Buffer
(FIFO)
Data
Consumer 1
resume/sleep
events
Data
Consumer 2
Variable event size / different events
Multiple producers/consumers
Selective consumer requests
Blocking/nonblocking requests
Robust: hanging client cannot crash system
High performance: many MB/sec throughput

## Slide 25: Event Structure
16-byte header
User data area free format
Optional MIDAS bank system
Optional YBOS bank system
Optional Subevent system
Default max. event size 512k
short int event_id
short int trg_mask
unsigned int serial_number
unsigned int time_stamp
unsigned int data_size
unsigned int data_size
unsigned int flags
bank_name[4]
type
size
data
bank_name[4]
type
size
data
subevent
subevent
subevent
subevent
subevent
subevent
subevent
subevent
16
Bytes
8
Bytes
8
Bytes
8
Bytes

## Slide 26: Event Requests
Cache
Consumer 1
event id, trigger mask
...
event id, trigger mask
Consumer 2
event id, trigger mask
...
Consumer 1
Consumer 2
FIFO
Buffer
Data
Producer
Consumer publishes specific event request(s)
Producer puts events in ring buffer and wakes up waiting consumers
If buffer is full, and consumer requested ALL events, producer is blocked (back-pressure)

## Slide 27: Remote Access to BM
Buffer
(FIFO)
Local Data
Consumer
Local Data
Producer
send_event()
IPC
spawn on
connect
MIDAS
server process
subprocess
send_event()
IPC
Remote Data
Producer
send_event()
Computer A
Computer B

## Slide 28: Multiple Experiments
Multiple independent experiments
Experiment name associated with directory on file system
Explicitly connect to experiment
FIFO
Online
FIFO
Offline
FIFO
Test
online /home/midas/online user1
offline /home/midas/offline user1
test /home/test test
/etc/exptab

## Slide 29: Watchdog Scheme
Data
Producer
events
Data
Consumer
Buffer
watchdog timer
watchdog timer
alarm()
update
check
alarm()
update
check

## Slide 30: Coding Conventions
Routine | Module | Meaning
cm_xxx | midas.c | Common routines (e.g. connect to experiment)
bm_xxx | midas.c | Buffer Manager routines
rpc_xxx | midas.c | Low-level RPC routines
bk_xxx | midas.c | Bank routines
ss_xxx | system.c | System Service routines
db_xxx | odb.c | Online Database routines
Data types:
“INT” 32-bit integer
“HNDLE” Integer number for object access (like a file handle)

## Slide 31: Simple producer & consumer
#include <stdio.h>
#include <stdlib.h>
#include "midas.h"
#define EVENT_SIZE 1000
main()
{
INT hBuf, status, start, stop, bytes_sent, serial_number=1;
char event[EVENT_SIZE], host_name[256];
printf("Host to connect: ");
ss_gets(host_name, 256);
/* connect to experiment */
cm_connect_experiment(host_name, "", "Producer", NULL);
/* open the event buffer with default size */
bm_open_buffer("SYSTEM", 0x100000, &hBuf);
/* set the buffer write cache size */
bm_set_cache_size(hBuf, 0, 50000);
do
{
start = ss_millitime();
bytes_sent = 0;
do
{
/* compose an event header with increasing serial number */
bm_compose_event((EVENT_HEADER *) event, 1, 1,
EVENT_SIZE, serial_number++);
/* send event */
rpc_send_event(hBuf, event, EVENT_SIZE+sizeof(EVENT_HEADER), SYNC);
bytes_sent += EVENT_SIZE;
/* repeat this loop for 1s */
} while (ss_millitime() - start < 1000);
/* Now calculate the rate */
stop = ss_millitime();
if (stop != start)
printf("Rate: %1.2lf MB/sec\n",
bytes_sent/1024.0/1024.0/(stop/1000.0 - start/1000.0));
status = cm_yield(0);
} while (status != RPC_SHUTDOWN && status != SS_ABORT);
cm_disconnect_experiment();
return 1;
}
#include <stdio.h>
#include <stdlib.h>
#include "midas.h"
void process_event(HNDLE hBuf, HNDLE request_id, EVENT_HEADER *pheader,
void *pevent)
{
static INT bytes_received, start=0, stop;
/* accumulate received event size */
bytes_received += pheader->data_size;
/* calculate rates each second */
if (ss_millitime() - start > 1000)
{
stop = ss_millitime();
if (stop != start)
printf("Rate: %1.2lf MB/sec\n",
bytes_received/1024.0/1024.0 /(stop/1000.0 - start/1000.0));
start = stop;
bytes_received = 0;
}
}
main()
{
INT hBuf, status, request_id;
/* connect to experiment */
cm_connect_experiment("", "", "Consumer", NULL);
/* open the event buffer */
bm_open_buffer("SYSTEM", 0x100000, &hBuf);
/* set the buffer cache size */
bm_set_cache_size(hBuf, 50000, 0);
/* place a request for a specific event id "1" */
bm_request_event(hBuf, 1, TRIGGER_ALL, GET_ALL, &request_id,
process_event);
do
{
status = cm_yield(1000);
} while (status != RPC_SHUTDOWN && status != SS_ABORT);
cm_disconnect_experiment();
return 1;
}
DEMO

## Slide 32: Performance
Locally (1.4 GHz Athlon, 1kB event size)
  160 MB/sec (one producer, one consumer)
  80 MB/sec (one producer, two consumers)
  5 MB/sec (10 Byte events)
100Mbit Ethernet
  10MB/sec
1Gbit Ethernet:
  limited by TCP/IP stack
  socket API: 30-50MB/sec
  zero-copy TCP/IP stack (send_file): under development, 2x possible

## Slide 33
Message System
Message
Producer
Buffer
(FIFO)
Alarm
Monitor
Alarm Message
Info Message
Message
Display
Messages are sent as events
Separate FIFO buffer
Message type = Event ID
Anyone can send or register for messages
Messages can be logged
File

## Slide 34: Online Database (ODB)

## Slide 35: Why database?
Store all variable experiment data
  Run number
  Logging channel information
  Event definition
  Front-end parameters (e.g. crate address)
  Analyzer parameters (e.g. pedestal values)
  Info about running programs
  Slow control values (HV, beam line)
Central database advantages
  “Single point of configuration”
  Avoid inconsistencies
  Easy backup
  Simple API for controlling the whole experiment

## Slide 36: Requirements
Fast access (>100.000/s), remote access
Tagged data: name, type, value
Hierachical structure (/proc, Windows registry)
Dynamical creation and deletion of keys
ASCII import/export
Automatic notification on data change
-> Requirements are not met by standard
(disk based) database

## Slide 37: Solution: ODB
Resides completely in shared memory
Access via MIDAS RPCs
Similar to file system
  File name -> “key”
  File contents -> key value (single value or array)
  Directory -> key list
  Symbolic link -> key link
  Access flags: read/write/delete
Data change notification via RPC callback
Persistency through disk image
Access through API and tools (ODBEdit)

## Slide 38
ODB Structure
db_set_value()
Client
Shared Memory
Runinfo
Run number
int
3
State
int
1
Logger
Channels
Data dir
char
/online/
0
Settings
Active
bool
1
Filename
char
run3.mid
.ODB.SHM
db_get_value()
Client
dump.odb
save/load
After last client stops
After run stop
db_flush_database()

## Slide 39: Corrupting the ODB
ODB has to be mapped into process address space
Misbehaving process (e.g. invalid pointer) can corrupt ODB
For debugging, all db_xxx routines can map/unmap address space
char *p;
db_set_value(…)
set_value
p = 12345;
*p = 0;
ODB
address space
char *p;
db_protect_database(1);
db_set_value(…)
ss_shmem_attach()
set_value
ss_shmem_detach()
p = 12345;
*p = 0;
Bad access immediately causes segmentation violation
Drawback:
Performance drop 400.000 acc/s -> 40.000 acc/s
Watchdog scheme doesn’t work anymore

## Slide 40: ODB ASCII format
[Experiment/Runinfo]
State = INT : 1
Online = BOOL : 1
Run number = DWORD : 10
Start time = STRING : [25] Wed Sep 04 10:03:47 1996
[Experiment/Runinfo/Flags]
Clear histos = BOOL : 1
Save histos = BOOL : 1
[Experiment/Analyzer parameters]
ADC pedestal = DOUBLE[4] :
1.2
4.3
2.5
5.6
Offset = DOUBLE : 9.5

## Slide 41: Data change notification (“hot link”)
subprocess A
db_set_value()
Remote
Client A
db_set_value()
Shared
Memory
run_number
subprocess B
notify_client()
notify flag B
Remote
Client B
callback()
db_open_record()
Computer B
Computer C
Computer A
4
1
2
3

## Slide 42: ODBEdit
Simple “wrapper” around MIDAS library
Old style command line interface
Inspect/change ODB, create new keys
Start/stop runs
Shutdown clients
Import/Export ODB to ASCII, C structure
Check ODB structure (“cleanup”)
Simple chat functionality
Change passwords
Works locally and remotely (-h <hostname>)

## Slide 43: Web interface
mhttpd
ODB interface
Run control
History display
Electronic Logbook
ODB
Web
Server
Client
Client
Client
RPC

## Slide 44: Example ODB Access
#include <stdio.h>
#include "midas.h"
int run_number;
void run_number_changed(HNDLE hDB, HNDLE hKey, void *info)
{
printf("Run number is %d\n", run_number);
}
main()
{
HNDLE hKey;
int size, status;
/* connect to experiment */
cm_connect_experiment("", "", "ODB Test", NULL);
/* get run number */
size = sizeof(run_number);
db_get_value(1, 0, "/runinfo/run number",
&run_number, &size, TID_INT, TRUE);
run_number++;
/* write run number */
db_set_value(1, 0, "/runinfo/run number",
&run_number, size, 1, TID_INT);
/* open hot link to run number */
db_find_key(1, 0, "/runinfo/run number", &hKey);
db_open_record(1, hKey, &run_number, sizeof(run_number),
MODE_READ, run_number_changed, NULL);
/* enter idle loop */
while (cm_yield(1000); != RPC_SHUTDOWN);
cm_disconnect_experiment();
return 1;
}
Callback routine gets called
whenever run number changes
Read, modify and write run_number
Register “hot link”
Event loop
odb1

## Slide 45: Hot link to C structure
struct {
int run_number;
int state;
} runinfo;
Runinfo
Run number
int
3
State
int
1
Application
ODB
-> Problem if ODB structure doesn’t match C structure
Solution:
correct ODB structure before hot-link is opened
lock ODB structure as long as hot-link is open

## Slide 46: Hot link to C structure example
typedef struct {
INT state;
INT online_mode;
INT run_number;
INT transition_in_progress;
INT requested_transition;
char start_time[32];
DWORD start_time_binary;
char stop_time[32];
DWORD stop_time_binary;
} RUN_INFO;
#define RUN_INFO_STR(_name) char *_name[] = {\
"[.]",\
"State = INT : 1",\
"Online Mode = INT : 1",\
"Run number = INT : 0",\
"Transition in progress = INT : 0",\
"Requested transition = INT : 0",\
"Start time = STRING : [32] Tue Sep 09 15:04:42 1997",\
"Start time binary = DWORD : 0",\
"Stop time = STRING : [32] Tue Sep 09 15:04:42 1997",\
"Stop time binary = DWORD : 0",\
"",\
NULL }
#include <stdio.h>
#include "midas.h"
#include “run_info.h”
RUN_INFO run_info;
RUN_INFO_STR(run_info_str);
void run_info_changed(HNDLE hDB, HNDLE hKey, void *info)
{
printf("Run number is %d\n", run_info.run_number);
}
main()
{
HNDLE hKey;
int status;
/* connect to experiment */
cm_connect_experiment("", "", "ODB Test", NULL);
/* create/correct /runinfo structure */
db_create_record(1, 0, "/Runinfo", strcomb(run_info_str));
/* open hot link to runinfo number */
db_find_key(1, 0, "/runinfo", &hKey);
db_open_record(1, hKey, &run_info, sizeof(run_info),
MODE_READ, run_info_changed, NULL);
/* enter idle loop */
do
{
status = cm_yield(1000);
} while (status != RPC_SHUTDOWN);
cm_disconnect_experiment();
return 1;
}
runinfo.h
Create ASCII representation of C structure
Use ASCII representation to create/correct ODB structure
odb2

## Slide 47: Examples for hot links
Steering of front-end:
struct {
int slot_adc;
int n_channels;
int read_adcs;
} conf;
…
if (conf.read_adcs)
for (i=0 ; i<conf.n_channels ; i++)
cami(0, conf.slot_adc, i, 0, adc[i]);
…
db_set_value(“/FE/Conf/Slot Adc”, 3);
struct {
int threshold1;
int threshold2;
} trigger_conf;
…
trigger_conf_changed(…)
{
camo(0,1,0,16, trigger_conf.threshold1);
camo(0,1,1,16, trigger_conf.threshold2);
}
Changing trigger hardware:
db_set_value(“/FE/Trigger Conf/Threshold1”, 1);

## Slide 48: Creating ASCII representations automatically with ODBEdit
[local:online:R]/>mkdir FE
[local:online:R]/>cd FE
[local:online:R]/FE>mkdir "Trigger Conf"
[local:online:R]/FE>cd "Trigger Conf"
[local:online:R]Trigger Conf>cre int Threshold1
[local:online:R]Trigger Conf>cre int Threshold2
[local:online:R]Trigger Conf>ls
Threshold1 0
Threshold2 0
[local:online:R]Trigger Conf>save -c conf.h
[local:online:R]Trigger Conf>save -s conf.c
typedef struct {
INT threshold1;
INT threshold2;
} TRIGGER_CONF;
#define TRIGGER_CONF(_name) char *_name[] = {\
"[.]",\
"Threshold1 = INT : 0",\
"Threshold2 = INT : 0",\
"",\
NULL }

## Slide 49: Slow Control System
ODB
HV demand
[0] 1000V
[1] 1500V
[2] 1300V
...
HV measured
[0] 1002V
[1] 1498V
[2] 1301V
...
HV Editor
Chn. HV
0 1002
1 1498
2 1301
1000
HV Frontend
HV power supply
RS232

## Slide 50
Run Transitions

## Slide 51: Run Transitions
Measurements are subdivided into “Runs”
Typically one Run is logged into one data file
Three run states: Stopped, Paused, Running
Run Transitions: Start, Stop, Pause, Resume
All clients need to be synchronized
running
paused
stopped
start
stop
resume
e.g. run start:
Logger opens data file
Analyzer clears histograms
Front-end enables trigger
pause

## Slide 52: Transition Synchronization
Remote
Client A
cm_transition()
transition
notify flag C
Remote
Client B
callback()
cm_register_transition()
transition
notify flag B
...
Client-Client RPC
ODB
1
2
3

## Slide 53: Transition example
#include <stdio.h>
#include "midas.h"
INT tr_start(INT run_number, char *error)
{
printf("Start run %d\n", run_number);
return CM_SUCCESS;
}
main()
{
int status;
cm_connect_experiment("", "", "Frontend", 0);
/* register to be notified on START transition */
cm_register_transition(TR_START, tr_start);
/* enter event loop */
while (cm_yield(1000) != RPC_SHUTDOWN);
cm_disconnect_experiment();
}
INT cm_transition(
INT transition,
INT run_number,
char *perror, INT strsize,
INT async_flag,
INT debug_flag);
transition
1
2
3

## Slide 54: Pre- Post- Transitions
Certain order necessary:
  e.g. first open logging file, then enable trigger
stopped
pre-start
start
post-start
open logging file
enable trigger
Only when all clients have successfully made one transition, continue with next

## Slide 55: MIDAS event loop
cm_yield(int timeout_ms)
check for events
bm_request_event()
check for messages
cm_msg_register()
check for transition
cm_register_transition()
check for key update
db_open_record()
check for RPC request
cm_register_function()
server socket
client socket
select(timeout)
select(timeout)
TCP/IP network
Register with:

## Slide 56
Slow Control System

## Slide 57: Slow Control
“Slow Control”: Everything not related to event-based readout
  High voltage
  Environment (temperature, pressure)
  Beam line magnets/slits
  Gas systems, …
Requirements
  Have all slow control data with event data on tape
  Access all slow control equipment in a standardized way
  Every program should have access to slow control equipment (analyzer, tools)
  History system (on-line display, date-to-date query)

## Slide 58: Slow Control through ODB
ODB
HV demand
HV Editor
1000.2
1000
HV Frontend
HV power
supply
HV measured
Frontend
Temperature
Pressure
Environment
variables
Hot-link
Temperature
Display
Beamline demand
Beamline measured
EPICS Frontend
TCP/IP
Analyzer
Copy image to datastream

## Slide 59: Three-layer driver concept
Class Driver
ODB
Device Driver
Bus Driver
connects to ODB
high level functionality (ramping, limits, …)
set channel
get channel
generate commands for specific device
read
write
send command to device
db_set_value(“/HV/Demand[5]”, 1000);
set_channel(5, 1000);
“WR5 V1000”
RS232
TCP/IP
GPIB
…

## Slide 60: Advantages
Class driver does not have to know anything about specific device
Device driver does not have to know how to physically connect to device
Bus driver does not have to know anything about device
Layers can be exchanged easily
Device driver very simple
New features in class driver are inherited by all devices automatically
Class driver can encapsulate slow control data in events for taping

## Slide 61: Slow Control Drivers
Available devices:
LeCroy HV
Keithley DAS cards
Epics Channel Access
Problem:
Devices are read out
sequentially, slow devices
may block others
-> “slowdev” class driver
-> separate front-ends
-> faster devices
midas/drivers/class/
midas/drivers/device/
midas/drivers/bus/
Front-end
Back-end
Front-end
Front-end
RS232
RS232
GPIB
PC
PC
PC

## Slide 62: Writing your own device driver
INT lrs1440_set(LRS1440_INFO *info, INT channel, float value)
{
char str[80];
sprintf(str, "W%04.0f C%02d\r", value, channel);
BD_PUTS(str);
BD_GETS(str, sizeof(str), "\r\n\r\n", 1000);
return FE_SUCCESS;
}
INT lrs1440_get(LRS1440_INFO *info, INT channel, float *pvalue)
{
int value;
char str[256];
sprintf(str, "R V C%02d\r", channel);
BD_PUTS(str);
BD_GETS(str, sizeof(str), "Act ", 1000);
BD_GETS(str, sizeof(str), "\r\n\r\n", 1000);
sscanf(str+1, "%d", &value);
*pvalue = (float) value;
return FE_SUCCESS;
}
Use existing driver as template
Implement device specific commands
Put device specific settings in C structure
Use BD_XXX routines for communication

## Slide 63: HVEdit
Windows/MFC
Qt
Web interface
Andreas Suter
PSI
C++ class library

## Slide 64: MSCB system
PC parallel
port
Bus powered
8-bit Microcontroller (Cygnal/Analog Devices)
2x12-bit DAC, 8x12-bit ADC
C programmable, Flash EEPROM
RS-485 @ 115kBaud
Up to 65536 nodes on 500m
Node cost: ~50 CN$
PSI designed cards for
  analog in 0-10V, 4-20mA
  digital in/out (24V)
  220V output
  Thermocouple
  PT100
  HV regulator(2400V/1mA)
output indicator LED
temperature sensors
READ
ADR 16bit
1 Byte
ACKN
VALUE
CRC
CHN
CRC
Protocol:
~1000 reads / sec.
DEMO

## Slide 65
History System

## Slide 66: History system
Requirements
  Want to store slow control data in database over years
  Online display of trend lines
  Database queries: “Show temperature last May”
  Fast write/search
Standard databases like MySQL are “overkill”
  no need for record locking, delete records, edit records
  CPU overhead
MIDAS history system
  small set of C functions
  events are linearly written to files, one file per day
  files contain event definition
  simple query function
  data gets only written when changing (hot-link) and when event is written to data stream

## Slide 67: History File Structure
event 1 definition
event 1
event 1
event 1
event 2 definition
event 2
event 1
event 2
event 2
event 1
event 2
event 2 definition
event 2
event 1
event 1 definition
event 1
event 1
event 2 definition
event 2
event 1
event 2
event 2
event 1
event 2
event 2
event 1
event 2
event 1
event 1 definition
event 1
event 1
event 2 definition
event 2
event 1
event 2
event 2
event 1
event 2
event 2
event 1
event 2
event 1
030516.hst
030517.hst
030518.hst
event definition can change during lifetime of experiment

## Slide 68: History Usage
History written by mlogger
Flag in front-end / slow control front-end
Link in ODB
Display settings (panels) in ODB
Key name Type #Val Size Last Opn Mode Value
---------------------------------------------------------------------------
History DIR
Links DIR
System DIR
Trigger per sec. LINK 1 46 1h 0 RWD /Equipment/Trigger/Statistics/Events per sec.
Trigger kBytes per sLINK 1 46 1h 0 RWD /Equipment/Trigger/Statistics/kBytes per sec.
Display DIR
Trigger rate DIR
Variables STRING 2 64 1h 0 RWD
[0] System:Trigger per sec.
[1] System:Trigger kB per sec.
Time Scale STRING 1 32 1h 0 RWD 1h
Factor FLOAT 2 4 1h 0 RWD
[0] 1
[1] 1
Offset FLOAT 2 4 1h 0 RWD
[0] 0
[1] 0
Timescale STRING 1 32 1h 0 RWD 1h
Zero ylow BOOL 1 4 1h 0 RWD y
Show run markers BOOL 1 4 1h 0 RWD y
Buttons STRING 7 32 1h 0 RWD
[0] 10m
[1] 1h
[2] 3h
[3] 12h
[4] 24h
[5] 3d
[6] 7d
Log axis BOOL 1 4 1h 0 RWD n

## Slide 69: Display configuration
Web based configuration
Resulting display
DEMO

## Slide 70:  Temperature Display
Display through mhttpd
Gif images generated dynamically in memory
mchart utility for strip-chart display
Query with mhist
Write speed: 2000 events/s
Query over one month in ~10sec

## Slide 71
Alarm System

## Slide 72: MIDAS Alarm System
Alarms are an important issue in experiments
Alarm System incorporated in MIDAS library -> every client does alarm checks
Four types of alarm
  ODB value <, >, =, !=
  Program stopped
  Periodic alarm (-> shift checks, change LN2)
  Internal alarms: Triggered by user code al_trigger_alarm()
Alarm triggers alarm “class” (Warning, Alarm, …)
  write system message (-> speaker)
  write elog message
  stop run
  execute command (-> pager, SMS)

## Slide 73: Alarm examples
Check value in ODB
Program stopped
Reset alarm
DEMO
Periodic alarm

## Slide 74
Electronic Logbook

## Slide 75: ELOG
Why give up paper logbook?
  Remote access (remote shifts, expert off-site)
  Easy copy
  Searchable
  Automatic logbook entries (analyzer)
How
  Web interface built into mhttpd (no CGI!)
  Simple flat file database (one file per day)
  Image attachments (scanner, digital camera)
  Standalone ELOG package (http://midas.psi.ch/elog)
Experience
  Excellent tool once experiment is in stable operation
  For setting up experiments, still some paper needed

## Slide 76: ELOG example
PAW screendump
run number automatically filled in

## Slide 77: ELOG forms
Form to be filled out every shift
Form items defined in ODB
Resulting ELOG entry

## Slide 78: Automatic ELOG entries
Modified HV’s
PMT ADC spectra have peak in spectrum
Spectra are cross-correlated with stored spectra, if enough statistics (~6h)
Gain deviation is calculated
HV is adjusted automatically (analyzer writes to ODB)
ELOG entry is produced

## Slide 79: Standalone ELOG
DEMO

## Slide 80
Frontend Framework

## Slide 81: “Simple” Front-end
#include <stdio.h>
#include "midas.h"
#include "msystem.h"
INT running; /* stopped:0, started:1 */
INT serial = 0; /* event serial number */
INT run_number = 0; /* run number */
HNDLE hBuf; /* buffer handle for events */
INT tr_start(INT rn, char *error)
{
printf("Start run %d\n", rn);
serial = 0;
run_number = rn;
running = TRUE;
return CM_SUCCESS;
}
INT tr_stop(INT rn, char *error)
{
running = FALSE;
rpc_flush_event();
bm_flush_cache(hBuf, SYNC);
printf("Stop run %d\n", rn);
return CM_SUCCESS;
}
main()
{
INT status, size=1000, id=0;
DWORD last_time=0;
char event[2000];
char host_name[256];
char dash[] = "-\\|/";
printf("Host to connect: ");
ss_gets(host_name, 256);
/* now connect to server */
status = cm_connect_experiment(host_name, "", "MiniFE", NULL);
if (status != CM_SUCCESS)
return 1;
cm_register_transition(TR_START, tr_start);
cm_register_transition(TR_STOP, tr_stop);
/* open event bufer and set cache size */
bm_open_buffer(EVENT_BUFFER_NAME, EVENT_BUFFER_SIZE, &hBuf);
bm_set_cache_size(hBuf, 0, 100000);
do
{
if (!running)
{
/* if run is not started, call yield with a timeout of 1s */
status = cm_yield(1000);
if (status == RPC_SHUTDOWN)
break;
}
else
{
/* during run call yield with 0 timeout once each 300 ms */
if (ss_millitime() - last_time > 300)
{
printf("%c\r", dash[id++ % 4]);
fflush(stdout);
status = cm_yield(0);
if (status == RPC_SHUTDOWN)
break;
last_time = ss_millitime();
}
}
/* if run is started, send event */
if (running)
{
/* compose event header, empty event */
bm_compose_event((EVENT_HEADER *) event, 1, 1, size, serial++);
/* send event */
rpc_send_event(hBuf, event, size+sizeof(EVENT_HEADER), SYNC);
}
} while (TRUE);
cm_disconnect_experiment();
return 1;
}
minife.c

## Slide 82: Front-end Framework
Update event statistics (events/s, kb/s)
Support for interrupt/polled/periodic/slow control events
Open buffer, send events
Send event samples to ODB for inspection
Stop run after number of events received
Incorporate slow control system
CNAF functionality
Trigger event declaration
Scaler event declaration
Scaler readout
Trigger event readout
BOR
EOR
Check for trigger event
Framework

## Slide 83: Front-end user code
#include <stdio.h>
#include <stdlib.h>
#include "midas.h"
#include "mcstd.h"
/*-- Equipment list --*/
EQUIPMENT equipment[] = {
{ "Trigger", /* equipment name */
{ 1, 0, /* event ID, trigger mask */
"SYSTEM", /* event buffer */
EQ_POLLED, /* equipment type */
LAM_SOURCE(0,0xFFFFFF),/* event source crate 0, all stations */
"MIDAS", /* format */
TRUE, /* enabled */
RO_RUNNING | /* read only when running */
RO_ODB, /* and update ODB */
500, /* poll for 500ms */
0, /* stop run after this event limit */
0, /* number of sub events */
0, /* don't log history */
"", "", "", },
read_trigger_event, /* readout routine */
NULL, NULL, NULL,
},
{ "Scaler", /* equipment name */
{ 2, 0, /* event ID, trigger mask */
"SYSTEM", /* event buffer */
EQ_PERIODIC |
EQ_MANUAL_TRIG, /* equipment type */
0, /* event source */
"MIDAS", /* format */
TRUE, /* enabled */
RO_RUNNING |
RO_TRANSITIONS | /* read when running and on transitions */
RO_ODB, /* and update ODB */
10000, /* read every 10 sec */
0, /* stop run after this event limit */
0, /* number of sub events */
1, /* log history */
"", "", "", },
read_scaler_event, /* readout routine */
NULL, NULL, NULL,
},
{ "" }
};
INT frontend_init() {
return SUCCESS;
}
INT frontend_exit() {
return SUCCESS;
}
INT begin_of_run(INT run_number, char *error) {
return SUCCESS;
}
INT end_of_run(INT run_number, char *error) {
return SUCCESS;
}
INT frontend_loop() {
return SUCCESS;
}
/*------------------------------------------------------------------*/
INT poll_event(INT source, INT count, BOOL test)
{
int i;
DWORD lam;
for (i=0 ; i<count ; i++)
{
cam_lam_read(LAM_SOURCE_CRATE(source), &lam);
if (lam & LAM_SOURCE_STATION(source))
if (!test)
return lam;
}
return 0;
}
INT read_trigger_event(char *pevent, INT off) {
return 0;
}
INT read_scaler_event(char *pevent, INT off) {
return 0;
}

## Slide 84: Creating banks
INT read_trigger_event(char *pevent, INT offset)
{
ADC0_BANK *adc0_bank;
WORD *pdata, a;
bk_init(pevent);
/* create structured ADC0 bank */
bk_create(pevent, "ADC0", TID_STRUCT, &adc0_bank);
pdata = (WORD *)adc0_bank;
/* read CAMAC adcs */
cami(CRATE, SLOT_ADC, 0, 0, &adc0_bank->adc0);
cami(CRATE, SLOT_ADC, 1, 0, &adc0_bank->adc1);
cami(CRATE, SLOT_ADC, 2, 0, &adc0_bank->adc2);
cami(CRATE, SLOT_ADC, 3, 0, &adc0_bank->adc3);
/* clear ADC */
camc(CRATE, SLOT_ADC, 0, 9);
bk_close(pevent, pdata);
/* create variable length TDC bank */
bk_create(pevent, "TDC0", TID_WORD, &pdata);
/* read CAMAC TDC until no Q */
cam16i_rq(CRATE, SLOT_TDC, 0, 0, &pdata, 96);
/* clear TDC */
camc(CRATE, SLOT_TDC, 0, 9);
bk_close(pevent, pdata);
return bk_size(pevent);
}
structured bank
typedef struct {
WORD adc0;
WORD adc1;
WORD adc2;
WORD adc3;
} ADC0_BANK;
bank with
variable length

## Slide 85: Frontend control through ODB
[local:online:S]/>cd Equipment/Trigger/
[local:online:S]Trigger>ls -lr
Key name Type #Val Size Last Opn Mode Value
---------------------------------------------------------------------------
Trigger DIR
Variables DIR
ADC0 DIR
adc0 WORD 1 2 1h 0 RWD 226
adc1 WORD 1 2 1h 0 RWD 29
adc2 WORD 1 2 1h 0 RWD 489
adc3 WORD 1 2 1h 0 RWD 574
TDC0 WORD 1 2 1h 0 RWD 0
[0] 49
[1] 152
[2] 156
[3] 223
ASUM DIR
Sum FLOAT 1 4 3h 0 RWD 1318
Average FLOAT 1 4 3h 0 RWD 329.5
Settings DIR
CAMAC slot BYTE 1 1 3h 0 RWD 7
Common DIR
Event ID WORD 1 2 1h 0 RWD 1
Trigger mask WORD 1 2 1h 0 RWD 0
Buffer STRING 1 32 1h 0 RWD SYSTEM
Type INT 1 4 1h 0 RWD 2
Source INT 1 4 1h 0 RWD 16777215
Format STRING 1 8 1h 0 RWD MIDAS
Enabled BOOL 1 4 1h 0 RWD y
Read on INT 1 4 1h 0 RWD 257
Period INT 1 4 1h 0 RWD 500
Event limit DOUBLE 1 8 1h 0 RWD 1e+006
Num subevents DWORD 1 4 1h 0 RWD 0
Log history INT 1 4 1h 0 RWD 0
Frontend host STRING 1 32 1h 0 RWD pc810
Frontend name STRING 1 32 1h 0 RWD Sample Frontend
Frontend file name STRING 1 256 1h 0 RWD c:\midas\examples\experiment\frontend.c
Statistics DIR
Events sent DOUBLE 1 8 0s 0 RWD 154234
Events per sec. DOUBLE 1 8 0s 0 RWD 132
kBytes per sec. DOUBLE 1 8 0s 0 RWD 12.233
structured
bank
Equipment
Structure
variable
length bank

## Slide 86: Hardware Access
Standardized hardware access layer makes front-end independent of interfaces
mcstd.h & esone.h
  cam8i(const int c, const int n, const int a, const int f, BYTE *d);
  cam16i_r(const int c, const int n, const int a, const int f, WORD **d, const int r);
  cam16i_sn(const int c, const int n, const int a, const int f, WORD **d, const int r);
mvmestd.h
  vme_open, vme_ioctl, vme_close, vme_read, vme_write, vme_mmap, vme_unmap
mfbstd.h
  fb_frd, fb_frc, fb_fwd, fb_fwc, …
vme_read(…)
User FE code
bt_read(…);
bt617.c
read(“/dev/vme”, …)
linux

## Slide 87: Equipment Types
Periodic
  scalers, …
Interrupt
  trigger events
  readout in interrupt routine
  dual buffer event transfer to front-end framework
Polled
  trigger events
  user-supplied polling routine
  readout routine can be debugged
  faster response time than interrupt on dedicated front-ends
Slowcontrol
  HV, environment, beam line, …
  Read by slow control system driver architecture

## Slide 88: Equipment Definition
#include "class/hv.h"
#include "device/lrs1440.h"
#include "device/lrs4032.h"
#include "bus/rs232.h"
/* device driver list */
DEVICE_DRIVER hv_driver[] = {
{ “LRS1440", lrs1440, 8, rs232 },
{ “LRS4032", lrs4032, 4, rs232 },
{ "" }
};
EQUIPMENT equipment[] = {
{ "HV", /* equipment name */
{ 3, 0, /* event ID, trigger mask */
"SYSTEM", /* event buffer */
EQ_SLOW, /* equipment type */
0, /* event source */
"FIXED", /* format */
TRUE, /* enabled */
RO_RUNNING | /* read when running */
RO_TRANSITIONS, /* and on transitions */
60000, /* read every 60 sec */
0, /* event limit */
0, /* number of sub events */
1, /* log history every event */
"", "", "", },
cd_hv_read, /* readout routine */
cd_hv, /* class driver main routine */
hv_driver, /* device driver list */
NULL, /* init string */
},
{ "" }
};
EQUIPMENT equipment[] = {
{ "Trigger", /* equipment name */
{ 1, 0, /* event ID, trigger mask */
"SYSTEM", /* event buffer */
EQ_POLLED, /* equipment type */
LAM_SOURCE(0,0xFFFFFF),/* event source crate 0, all stations */
"MIDAS", /* format */
TRUE, /* enabled */
RO_RUNNING | /* read only when running */
RO_ODB, /* and update ODB */
500, /* poll for 500ms */
0, /* stop run after this event limit */
0, /* number of sub events */
0, /* don't log history */
"", "", "", },
read_trigger_event, /* readout routine */
NULL, NULL, NULL,
},
{ "Scaler", /* equipment name */
{ 2, 0, /* event ID, trigger mask */
"SYSTEM", /* event buffer */
EQ_PERIODIC, /* equipment type */
0, /* event source */
"MIDAS", /* format */
TRUE, /* enabled */
RO_RUNNING |
RO_TRANSITIONS | /* read when running and on transitions */
RO_ODB, /* and update ODB */
10000, /* read every 10 sec */
0, /* stop run after this event limit */
0, /* number of sub events */
0, /* log history */
"", "", "", },
read_scaler_event, /* readout routine */
NULL, NULL, NULL,
},
};
Polled, Periodic:
Slowcontrol:

## Slide 89
Analyzer Framework

## Slide 90: MIDAS analyzer framework (mana.c)
Automatically receive events
  Online (buffer manager), request only certain events
  Offline (read from *.mid file)
Write statistics to ODB
Generate histos
  Online
    HBOOK histos and online N-tuples (shared memory)
    ROOT histos (socket server)
  Offline
    Write *.rz file with RWNT or CWNT and histos
    Write *.root file with TTree and histos
Modular concept
Get parameters (e.g. calibration constants) from ODB

## Slide 91: Modular concept
Online data arrives in banks
Events get extended with “calculated” banks by analysis module
Each module is independent of others
Typically first calibration modules, finally “physics” modules calculating particle properties
Collaboration agrees on common low-level calibration module
High-level modules get personalized
Per-bank switch for N-tuple creation
Each module has its own parameters (mapped to ODB)
Parameters can be changed offline

## Slide 92: Analyzer Example
ANA_MODULE *trigger_module[] = {
&adc_calib_module,
&adc_summing_module,
NULL
};
/*-- Bank definitions --*/
ASUM_BANK_STR(asum_bank_str);
ADC0_BANK_STR(ana_adc0_bank_str);
BANK_LIST ana_trigger_bank_list[] = {
/* online banks */
{ "ADC0", TID_STRUCT, sizeof(ADC0_BANK), ana_adc0_bank_str },
{ "TDC0", TID_WORD, N_TDC, NULL },
/* calculated banks */
{ "CADC", TID_FLOAT, N_ADC, NULL },
{ "ASUM", TID_STRUCT, sizeof(ASUM_BANK), asum_bank_str },
{ "" },
};
/*-- Event request list --*/
ANALYZE_REQUEST analyze_request[] = {
{ "Trigger", /* equipment name */
{ 1, /* event ID */
TRIGGER_ALL, /* trigger mask */
GET_SOME, /* get some events */
"SYSTEM", /* event buffer */
TRUE, /* enabled */
"", "", },
NULL, /* analyzer routine */
trigger_module, /* module list */
ana_trigger_bank_list,/* bank list */
1000, /* RWNT buffer size */
TRUE, /* Use tests for this event */
},
};
ADC_SUMMING_PARAM adc_summing_param;
ADC_SUMMING_PARAM_STR(adc_summing_param_str);
ANA_MODULE adc_summing_module = {
"ADC summing", /* module name */
adc_summing, /* event routine */
adc_summing_init, /* init routine */
&adc_summing_param, /* parameter structure */
sizeof(adc_summing_param), /* structure size */
adc_summing_param_str, /* initial parameters */
};
static TH1F *gAdcSumHist;
INT adc_summing_init(void) {
/* book ADCSUM histo */
gAdcSumHist = new TH1F("ADCSUM", "ADC sum", 500, 0, 10000);
return SUCCESS;
}
INT adc_summing(EVENT_HEADER *pheader, void *pevent) {
INT i, j, n_adc;
float *cadc;
ASUM_BANK *asum;
/* locate CADC bank, create ASUM bank */
n_adc = bk_locate(pevent, "CADC", &cadc);
bk_create(pevent, "ASUM", TID_STRUCT, &asum);
/* sum all channels above threashold */
asum->sum = 0.f;
for (i=j=0 ; i<n_adc ; i++)
if (cadc[i] > adc_summing_param.adc_threshold) {
asum->sum += cadc[i];
j++;
}
/* calculate ADC average */
asum->average = j > 0 ? asum->sum / j : 0;
/* fill sum histo */
gAdcSumHist->Fill(asum->sum, 1);
/* close calculated bank */
bk_close(pevent, asum+1);
return SUCCESS;
}
analyzer.c
adcsum.c

## Slide 93: Offline analysis
CADC
TDC
ADC
TDC
PCOS
run123.mid
analyzer
ODB
CADC
MWPC
ADC
TDC
PCOS
run123.rz
Analyzer read online banks
and adds calculated data
selected banks are
converted to N-tuples
and written to output file
PAW
ROOT
analyzer –i run123.mid –o run123.rz
ODB dump
run123.root

## Slide 94: Calibration constants
analyzer
ODB
/analyzer/parameters/ADC summing
ADC threshold = 5
…
run123.mid
mlogger
Online
Offline
analyzer
ODB
analyzer
calib.odb
analyzer –i run123.mid –o run123.rz –c calib.odb
[/Analyzer/Parameters/ADC summing]
ADC threshold = FLOAT : 3
original online constants get
loaded from *.mid file but
can be overloaded

## Slide 95: ODB sections for analyzer
/Equipment/<name>/Variables | Contains the event definition of a specific event. [name] can be "trigger", "scaler", "hv" and so on.
/Runinfo | Contains information about the current run.
/Analyzer/Module Switches/<module> | Enable switches for all modules. A "0" switches a module off during analysis.
/Analyzer/Parameters/<module> | Contains parameters for all analyzer modules.
/Analyzer/Bank Switches | Output flags for all banks. A "1" indicates that the given bank should be included in the analyzer output.
/Analyzer/Output | Info about the analyzer output.

## Slide 96: C structure – ODB mapping
Frontend
Analyzer
typedef struct {
WORD adc0;
WORD adc1;
WORD adc2;
WORD adc3;
} ADC0_BANK;
/Equipment/Trigger/Variables/
Key name Type #Val Size Last Opn Mode Value
---------------------------------------------------------------------------
Variables DIR
ADC0 DIR
adc0 WORD 1 2 22h 0 RWD 881
adc1 WORD 1 2 22h 0 RWD 152
adc2 WORD 1 2 22h 0 RWD 762
adc3 WORD 1 2 22h 0 RWD 643
ASUM DIR
Sum FLOAT 1 4 4h 0 RWD 425.5
Average FLOAT 1 4 4h 0 RWD 141.833
typedef struct {
float sum;
float average;
} ADC0_BANK;
/Equipment/Trigger/Settings/
Key name Type #Val Size Last Opn Mode Value
---------------------------------------------------------------------------
ADC slot WORD 1 2 22h 0 RWD 1
TDC slot WORD 1 2 22h 0 RWD 2
typedef struct {
WORD adc_slot;
WORD tdc_slot;
} TRIGGER_SETTINGS;
/Analyzer/Parameters/ADC summing/
Key name Type #Val Size Last Opn Mode Value
---------------------------------------------------------------------------
ADC threshold FLOAT 1 2 22h 0 RWD 10
Pedestal FLOAT 1 2 22h 0 RWD -3.5
typedef struct {
float adc threshold;
float pedestal;
} ADC_SUMMING_PARAM;
create ODB entries
enter “make” in ODBEdit
“experim.h” is automatically created
see also:
http://pibeta.web.psi.ch/handbook/analyzer/

## Slide 97: FAL
For simple experiment, front-end + analyzer + logger are combined in single process
Avoid overhead of inter-process communication
Only single process needs to be started
Disadvantage: crashing analyzer stops DAQ
front-end
analyzer
logger
PAW
run123.mid
hardware

## Slide 98: “Histogram” readout
Event-based readout
  ADC/TDC read by frontend
  polled/interrupt events
  analyzer does histogram filling in software
Histogram readout
  Histograms filled in hardware
  periodic events
  analyzer displays histos
ADC
TDC
FE
Analyzer
ADC
TDC
FE
Analyzer
adc/tdc
events
histo
events
104 events/s
106 events/s
<1 event/s

## Slide 99: Online Demo
Get tar file from http://midas.psi.ch/
Unpack tar file tar xzvf midas-x.tar.gz
Compile MIDAS make
Install MIDAS make install → /usr/local
Define Experiment(s) /etc/exptab
Enable mserver /etc/services /etc/xinetd.d/midas
Compile experiment examples/experiment> make
Run experiment frontend analyzer mlogger mhttpd rmidas

## Slide 100: Part II
Online demo
  Installation
  Configuration
  Running a small experiment
Utilities and usage
Known deployment
Advanced features
  Customization of run control and monitoring
  Special front-end features
Advanced online demo